## T150263
SN: | T150263
Item #: | TW930R5A2
Item Description: | 2x2 50:50 Wideband Fiber Optic Coupler
Fiber Type: | Thorlabs 780HP
Port Name: | White (Input) - White (Signal Output)
Test Wavelength (nm): | 930
Insertion Loss (dB): | -3.29
Coupling Ratio (%): | 47.46
Uniformity (dB): | 0.52
Port Name: | White (Input) - Red (Tap Output)
Test Wavelength (nm): | 930
Insertion Loss (dB): | -2.84
Coupling Ratio (%): | 52.54
Uniformity (dB): | 0.5
Spectral Response
Wavelength (nm) | Insertion Loss - White (Signal Output) (dB) | Coupling Ratio - White (Signal Output) (%) | Insertion Loss - Red (Tap Output) (dB) | Coupling Ratio - Red (Tap Output) (%)
820 | -2.751 | 53.572 | -3.373 | 46.428
820.441 | -2.75 | 53.557 | -3.369 | 46.443
820.882 | -2.749 | 53.531 | -3.363 | 46.469
821.323 | -2.759 | 53.46 | -3.361 | 46.54
821.764 | -2.773 | 53.375 | -3.36 | 46.625
822.204 | -2.77 | 53.359 | -3.354 | 46.641
822.645 | -2.767 | 53.34 | -3.348 | 46.66
823.086 | -2.773 | 53.288 | -3.345 | 46.712
823.527 | -2.779 | 53.243 | -3.343 | 46.757
823.968 | -2.785 | 53.211 | -3.344 | 46.789
824.409 | -2.79 | 53.17 | -3.341 | 46.83
824.85 | -2.792 | 53.125 | -3.336 | 46.875
825.291 | -2.794 | 53.104 | -3.334 | 46.896
825.731 | -2.796 | 53.088 | -3.333 | 46.912
826.172 | -2.801 | 53.048 | -3.331 | 46.952
826.613 | -2.806 | 52.996 | -3.327 | 47.004
827.054 | -2.811 | 52.9 | -3.315 | 47.1
827.495 | -2.817 | 52.826 | -3.308 | 47.174
827.936 | -2.823 | 52.785 | -3.307 | 47.215
828.377 | -2.826 | 52.74 | -3.303 | 47.26
828.818 | -2.828 | 52.69 | -3.296 | 47.31
829.259 | -2.833 | 52.609 | -3.286 | 47.391
829.699 | -2.838 | 52.525 | -3.277 | 47.475
830.14 | -2.846 | 52.506 | -3.282 | 47.494
830.581 | -2.851 | 52.494 | -3.285 | 47.506
831.022 | -2.848 | 52.496 | -3.282 | 47.504
831.463 | -2.85 | 52.473 | -3.28 | 47.527
831.904 | -2.857 | 52.421 | -3.277 | 47.579
832.345 | -2.864 | 52.368 | -3.276 | 47.632
832.786 | -2.871 | 52.313 | -3.273 | 47.687
833.226 | -2.875 | 52.238 | -3.264 | 47.762
833.667 | -2.878 | 52.163 | -3.254 | 47.837
834.108 | -2.881 | 52.123 | -3.25 | 47.877
834.549 | -2.884 | 52.105 | -3.25 | 47.895
834.99 | -2.887 | 52.125 | -3.257 | 47.875
835.431 | -2.891 | 52.091 | -3.254 | 47.909
835.872 | -2.895 | 52.001 | -3.243 | 47.999
836.313 | -2.898 | 51.967 | -3.24 | 48.033
836.753 | -2.901 | 51.958 | -3.241 | 48.042
837.194 | -2.905 | 51.9 | -3.235 | 48.1
837.635 | -2.91 | 51.836 | -3.229 | 48.164
838.076 | -2.915 | 51.789 | -3.226 | 48.211
838.517 | -2.918 | 51.706 | -3.214 | 48.294
838.958 | -2.918 | 51.555 | -3.188 | 48.445
839.399 | -2.92 | 51.439 | -3.17 | 48.561
839.84 | -2.925 | 51.365 | -3.162 | 48.635
840.281 | -2.929 | 51.311 | -3.157 | 48.689
840.721 | -2.934 | 51.256 | -3.153 | 48.744
841.162 | -2.94 | 51.13 | -3.136 | 48.87
841.603 | -2.945 | 51.019 | -3.122 | 48.981
842.044 | -2.952 | 51.008 | -3.127 | 48.992
842.485 | -2.955 | 51.016 | -3.131 | 48.984
842.926 | -2.953 | 51.032 | -3.132 | 48.968
843.367 | -2.955 | 51 | -3.129 | 49
843.808 | -2.96 | 50.932 | -3.122 | 49.068
844.248 | -2.962 | 50.876 | -3.114 | 49.124
844.689 | -2.962 | 50.822 | -3.105 | 49.178
845.13 | -2.973 | 50.727 | -3.099 | 49.273
845.571 | -2.982 | 50.651 | -3.095 | 49.349
846.012 | -2.986 | 50.648 | -3.098 | 49.352
846.453 | -2.991 | 50.639 | -3.102 | 49.361
846.894 | -3 | 50.61 | -3.106 | 49.39
847.335 | -3.005 | 50.579 | -3.106 | 49.421
847.776 | -3.008 | 50.551 | -3.104 | 49.449
848.216 | -3.012 | 50.532 | -3.104 | 49.468
848.657 | -3.016 | 50.516 | -3.105 | 49.484
849.098 | -3.018 | 50.486 | -3.102 | 49.514
849.539 | -3.019 | 50.451 | -3.098 | 49.549
849.98 | -3.019 | 50.407 | -3.09 | 49.593
850.421 | -3.022 | 50.366 | -3.086 | 49.634
850.862 | -3.028 | 50.332 | -3.086 | 49.668
851.303 | -3.034 | 50.309 | -3.087 | 49.691
851.743 | -3.038 | 50.289 | -3.088 | 49.711
852.184 | -3.039 | 50.257 | -3.083 | 49.743
852.625 | -3.038 | 50.222 | -3.077 | 49.778
853.066 | -3.041 | 50.193 | -3.074 | 49.807
853.507 | -3.045 | 50.155 | -3.072 | 49.845
853.948 | -3.052 | 50.1 | -3.069 | 49.9
854.389 | -3.057 | 50.054 | -3.067 | 49.946
854.83 | -3.06 | 50.022 | -3.064 | 49.978
855.271 | -3.064 | 49.994 | -3.063 | 50.006
855.711 | -3.066 | 49.975 | -3.061 | 50.025
856.152 | -3.062 | 49.99 | -3.061 | 50.01
856.593 | -3.058 | 50.005 | -3.059 | 49.995
857.034 | -3.056 | 49.977 | -3.052 | 50.023
857.475 | -3.056 | 49.94 | -3.046 | 50.06
857.916 | -3.057 | 49.905 | -3.041 | 50.095
858.357 | -3.057 | 49.896 | -3.039 | 50.104
858.798 | -3.055 | 49.905 | -3.039 | 50.095
859.238 | -3.059 | 49.864 | -3.036 | 50.136
859.679 | -3.066 | 49.793 | -3.031 | 50.207
860.12 | -3.073 | 49.723 | -3.024 | 50.277
860.561 | -3.076 | 49.668 | -3.019 | 50.332
861.002 | -3.075 | 49.655 | -3.015 | 50.345
861.443 | -3.076 | 49.637 | -3.013 | 50.363
861.884 | -3.081 | 49.6 | -3.012 | 50.4
862.325 | -3.087 | 49.568 | -3.012 | 50.432
862.766 | -3.091 | 49.544 | -3.012 | 50.456
863.206 | -3.091 | 49.528 | -3.009 | 50.472
863.647 | -3.09 | 49.514 | -3.005 | 50.486
864.088 | -3.09 | 49.49 | -3.001 | 50.51
864.529 | -3.09 | 49.466 | -2.997 | 50.534
864.97 | -3.091 | 49.446 | -2.995 | 50.554
865.411 | -3.092 | 49.424 | -2.992 | 50.576
865.852 | -3.093 | 49.402 | -2.99 | 50.598
866.293 | -3.094 | 49.394 | -2.989 | 50.606
866.733 | -3.095 | 49.397 | -2.99 | 50.603
867.174 | -3.097 | 49.397 | -2.993 | 50.603
867.615 | -3.101 | 49.392 | -2.995 | 50.608
868.056 | -3.105 | 49.376 | -2.996 | 50.624
868.497 | -3.109 | 49.354 | -2.997 | 50.646
868.938 | -3.113 | 49.324 | -2.995 | 50.676
869.379 | -3.115 | 49.294 | -2.993 | 50.706
869.82 | -3.116 | 49.269 | -2.989 | 50.731
870.26 | -3.114 | 49.25 | -2.984 | 50.75
870.701 | -3.111 | 49.232 | -2.978 | 50.768
871.142 | -3.109 | 49.201 | -2.97 | 50.799
871.583 | -3.109 | 49.161 | -2.963 | 50.839
872.024 | -3.112 | 49.113 | -2.958 | 50.887
872.465 | -3.115 | 49.069 | -2.953 | 50.931
872.906 | -3.116 | 49.038 | -2.949 | 50.962
873.347 | -3.117 | 49.019 | -2.947 | 50.981
873.788 | -3.119 | 49.009 | -2.947 | 50.991
874.228 | -3.121 | 48.993 | -2.946 | 51.007
874.669 | -3.123 | 48.976 | -2.946 | 51.024
875.11 | -3.124 | 48.976 | -2.946 | 51.024
875.551 | -3.125 | 48.989 | -2.949 | 51.011
875.992 | -3.126 | 49.009 | -2.953 | 50.991
876.433 | -3.126 | 49.019 | -2.956 | 50.981
876.874 | -3.127 | 49.003 | -2.954 | 50.997
877.315 | -3.129 | 48.974 | -2.951 | 51.026
877.755 | -3.134 | 48.941 | -2.95 | 51.059
878.196 | -3.139 | 48.923 | -2.952 | 51.077
878.637 | -3.143 | 48.915 | -2.954 | 51.085
879.078 | -3.143 | 48.91 | -2.953 | 51.09
879.519 | -3.141 | 48.903 | -2.951 | 51.097
879.96 | -3.141 | 48.886 | -2.947 | 51.114
880.401 | -3.143 | 48.852 | -2.944 | 51.148
880.842 | -3.149 | 48.805 | -2.942 | 51.195
881.283 | -3.153 | 48.786 | -2.942 | 51.214
881.723 | -3.154 | 48.789 | -2.943 | 51.211
882.164 | -3.154 | 48.763 | -2.939 | 51.237
882.605 | -3.155 | 48.719 | -2.932 | 51.281
883.046 | -3.157 | 48.69 | -2.929 | 51.31
883.487 | -3.158 | 48.668 | -2.927 | 51.332
883.928 | -3.159 | 48.644 | -2.923 | 51.356
884.369 | -3.161 | 48.622 | -2.921 | 51.378
884.81 | -3.165 | 48.602 | -2.922 | 51.398
885.25 | -3.171 | 48.574 | -2.924 | 51.426
885.691 | -3.178 | 48.535 | -2.924 | 51.465
886.132 | -3.184 | 48.48 | -2.92 | 51.52
886.573 | -3.189 | 48.432 | -2.916 | 51.568
887.014 | -3.191 | 48.431 | -2.918 | 51.569
887.455 | -3.191 | 48.444 | -2.921 | 51.556
887.896 | -3.191 | 48.445 | -2.92 | 51.555
888.337 | -3.19 | 48.429 | -2.917 | 51.571
888.778 | -3.191 | 48.396 | -2.912 | 51.604
889.218 | -3.193 | 48.354 | -2.907 | 51.646
889.659 | -3.195 | 48.318 | -2.902 | 51.682
890.1 | -3.195 | 48.309 | -2.901 | 51.691
890.541 | -3.197 | 48.309 | -2.904 | 51.691
890.982 | -3.206 | 48.292 | -2.909 | 51.708
891.423 | -3.212 | 48.275 | -2.913 | 51.725
891.864 | -3.212 | 48.269 | -2.912 | 51.731
892.305 | -3.212 | 48.259 | -2.909 | 51.741
892.745 | -3.213 | 48.232 | -2.906 | 51.768
893.186 | -3.214 | 48.187 | -2.899 | 51.813
893.627 | -3.215 | 48.141 | -2.892 | 51.859
894.068 | -3.218 | 48.123 | -2.892 | 51.877
894.509 | -3.223 | 48.12 | -2.896 | 51.88
894.95 | -3.227 | 48.113 | -2.899 | 51.887
895.391 | -3.23 | 48.095 | -2.899 | 51.905
895.832 | -3.232 | 48.056 | -2.894 | 51.944
896.273 | -3.234 | 47.992 | -2.885 | 52.008
896.713 | -3.238 | 47.92 | -2.876 | 52.08
897.154 | -3.242 | 47.882 | -2.873 | 52.118
897.595 | -3.245 | 47.862 | -2.873 | 52.138
898.036 | -3.249 | 47.813 | -2.869 | 52.187
898.477 | -3.251 | 47.778 | -2.865 | 52.222
898.918 | -3.252 | 47.806 | -2.87 | 52.194
899.359 | -3.251 | 47.844 | -2.876 | 52.156
899.8 | -3.251 | 47.85 | -2.877 | 52.15
900.24 | -3.253 | 47.826 | -2.875 | 52.174
900.681 | -3.257 | 47.787 | -2.872 | 52.213
901.122 | -3.262 | 47.764 | -2.873 | 52.236
901.563 | -3.266 | 47.751 | -2.875 | 52.249
902.004 | -3.269 | 47.728 | -2.874 | 52.272
902.445 | -3.27 | 47.7 | -2.87 | 52.3
902.886 | -3.268 | 47.678 | -2.864 | 52.322
903.327 | -3.266 | 47.665 | -2.86 | 52.335
903.768 | -3.265 | 47.656 | -2.858 | 52.344
904.208 | -3.267 | 47.644 | -2.857 | 52.356
904.649 | -3.269 | 47.632 | -2.857 | 52.368
905.09 | -3.27 | 47.627 | -2.858 | 52.373
905.531 | -3.273 | 47.621 | -2.86 | 52.379
905.972 | -3.28 | 47.596 | -2.862 | 52.404
906.413 | -3.285 | 47.57 | -2.863 | 52.43
906.854 | -3.286 | 47.561 | -2.862 | 52.439
907.295 | -3.287 | 47.556 | -2.862 | 52.444
907.735 | -3.289 | 47.551 | -2.863 | 52.449
908.176 | -3.289 | 47.553 | -2.863 | 52.447
908.617 | -3.287 | 47.557 | -2.862 | 52.443
909.058 | -3.289 | 47.553 | -2.863 | 52.447
909.499 | -3.293 | 47.527 | -2.863 | 52.473
909.94 | -3.297 | 47.475 | -2.858 | 52.525
910.381 | -3.298 | 47.436 | -2.853 | 52.564
910.822 | -3.297 | 47.436 | -2.851 | 52.564
911.263 | -3.295 | 47.43 | -2.848 | 52.57
911.703 | -3.293 | 47.409 | -2.843 | 52.591
912.144 | -3.29 | 47.409 | -2.84 | 52.591
912.585 | -3.287 | 47.428 | -2.84 | 52.572
913.026 | -3.289 | 47.431 | -2.843 | 52.569
913.467 | -3.294 | 47.419 | -2.845 | 52.581
913.908 | -3.299 | 47.405 | -2.848 | 52.595
914.349 | -3.301 | 47.398 | -2.849 | 52.602
914.79 | -3.3 | 47.401 | -2.848 | 52.599
915.23 | -3.295 | 47.411 | -2.845 | 52.589
915.671 | -3.29 | 47.421 | -2.841 | 52.579
916.112 | -3.288 | 47.42 | -2.839 | 52.58
916.553 | -3.288 | 47.412 | -2.838 | 52.588
916.994 | -3.29 | 47.4 | -2.838 | 52.6
917.435 | -3.292 | 47.375 | -2.835 | 52.625
917.876 | -3.294 | 47.328 | -2.83 | 52.672
918.317 | -3.294 | 47.306 | -2.826 | 52.694
918.758 | -3.292 | 47.329 | -2.828 | 52.671
919.198 | -3.291 | 47.349 | -2.83 | 52.651
919.639 | -3.292 | 47.344 | -2.83 | 52.656
920.08 | -3.293 | 47.341 | -2.831 | 52.659
920.521 | -3.294 | 47.343 | -2.831 | 52.657
920.962 | -3.293 | 47.339 | -2.83 | 52.661
921.403 | -3.292 | 47.338 | -2.829 | 52.662
921.844 | -3.29 | 47.35 | -2.829 | 52.65
922.285 | -3.289 | 47.358 | -2.83 | 52.642
922.725 | -3.29 | 47.353 | -2.83 | 52.647
923.166 | -3.294 | 47.351 | -2.833 | 52.649
923.607 | -3.298 | 47.354 | -2.838 | 52.646
924.048 | -3.302 | 47.355 | -2.843 | 52.645
924.489 | -3.304 | 47.355 | -2.844 | 52.645
924.93 | -3.302 | 47.359 | -2.843 | 52.641
925.371 | -3.297 | 47.384 | -2.842 | 52.616
925.812 | -3.291 | 47.435 | -2.845 | 52.565
926.253 | -3.287 | 47.481 | -2.849 | 52.519
926.693 | -3.287 | 47.495 | -2.851 | 52.505
927.134 | -3.288 | 47.485 | -2.851 | 52.515
927.575 | -3.29 | 47.469 | -2.85 | 52.531
928.016 | -3.291 | 47.465 | -2.85 | 52.535
928.457 | -3.291 | 47.471 | -2.851 | 52.529
928.898 | -3.29 | 47.473 | -2.851 | 52.527
929.339 | -3.289 | 47.47 | -2.849 | 52.53
929.78 | -3.287 | 47.463 | -2.846 | 52.537
930.22 | -3.285 | 47.462 | -2.844 | 52.538
930.661 | -3.285 | 47.476 | -2.846 | 52.524
931.102 | -3.286 | 47.504 | -2.852 | 52.496
931.543 | -3.287 | 47.536 | -2.859 | 52.464
931.984 | -3.289 | 47.559 | -2.865 | 52.441
932.425 | -3.291 | 47.562 | -2.868 | 52.438
932.866 | -3.293 | 47.543 | -2.866 | 52.457
933.307 | -3.294 | 47.525 | -2.864 | 52.475
933.747 | -3.293 | 47.524 | -2.863 | 52.476
934.188 | -3.291 | 47.534 | -2.863 | 52.466
934.629 | -3.289 | 47.54 | -2.862 | 52.46
935.07 | -3.288 | 47.528 | -2.858 | 52.472
935.511 | -3.287 | 47.505 | -2.853 | 52.495
935.952 | -3.287 | 47.492 | -2.851 | 52.508
936.393 | -3.286 | 47.492 | -2.85 | 52.508
936.834 | -3.285 | 47.503 | -2.851 | 52.497
937.275 | -3.285 | 47.514 | -2.852 | 52.486
937.715 | -3.284 | 47.524 | -2.854 | 52.476
938.156 | -3.283 | 47.542 | -2.855 | 52.458
938.597 | -3.28 | 47.569 | -2.857 | 52.431
939.038 | -3.278 | 47.585 | -2.858 | 52.415
939.479 | -3.278 | 47.588 | -2.858 | 52.412
939.92 | -3.278 | 47.584 | -2.858 | 52.416
940.361 | -3.277 | 47.595 | -2.858 | 52.405
940.802 | -3.27 | 47.636 | -2.86 | 52.364
941.242 | -3.265 | 47.682 | -2.862 | 52.318
941.683 | -3.263 | 47.712 | -2.866 | 52.288
942.124 | -3.264 | 47.734 | -2.87 | 52.266
942.565 | -3.264 | 47.753 | -2.873 | 52.247
943.006 | -3.263 | 47.77 | -2.875 | 52.23
943.447 | -3.261 | 47.782 | -2.875 | 52.218
943.888 | -3.259 | 47.785 | -2.874 | 52.215
944.329 | -3.258 | 47.789 | -2.873 | 52.211
944.77 | -3.255 | 47.803 | -2.873 | 52.197
945.21 | -3.253 | 47.817 | -2.874 | 52.183
945.651 | -3.253 | 47.825 | -2.875 | 52.175
946.092 | -3.253 | 47.829 | -2.875 | 52.171
946.533 | -3.252 | 47.834 | -2.876 | 52.166
946.974 | -3.252 | 47.842 | -2.877 | 52.158
947.415 | -3.25 | 47.85 | -2.877 | 52.15
947.856 | -3.249 | 47.85 | -2.875 | 52.15
948.297 | -3.247 | 47.842 | -2.872 | 52.158
948.737 | -3.246 | 47.83 | -2.869 | 52.17
949.178 | -3.246 | 47.834 | -2.87 | 52.166
949.619 | -3.246 | 47.858 | -2.874 | 52.142
950.06 | -3.245 | 47.883 | -2.877 | 52.117
950.501 | -3.242 | 47.902 | -2.877 | 52.098
950.942 | -3.239 | 47.927 | -2.878 | 52.073
951.383 | -3.236 | 47.962 | -2.882 | 52.038
951.824 | -3.235 | 48.005 | -2.888 | 51.995
952.265 | -3.233 | 48.044 | -2.893 | 51.956
952.705 | -3.231 | 48.072 | -2.896 | 51.928
953.146 | -3.228 | 48.093 | -2.897 | 51.907
953.587 | -3.225 | 48.112 | -2.897 | 51.888
954.028 | -3.225 | 48.124 | -2.899 | 51.876
954.469 | -3.226 | 48.132 | -2.901 | 51.868
954.91 | -3.223 | 48.139 | -2.899 | 51.861
955.351 | -3.218 | 48.139 | -2.895 | 51.861
955.792 | -3.216 | 48.123 | -2.89 | 51.877
956.232 | -3.215 | 48.12 | -2.889 | 51.88
956.673 | -3.215 | 48.148 | -2.893 | 51.852
957.114 | -3.215 | 48.175 | -2.898 | 51.825
957.555 | -3.215 | 48.181 | -2.899 | 51.819
957.996 | -3.214 | 48.168 | -2.896 | 51.832
958.437 | -3.213 | 48.145 | -2.891 | 51.855
958.878 | -3.211 | 48.123 | -2.885 | 51.877
959.319 | -3.209 | 48.113 | -2.881 | 51.887
959.76 | -3.212 | 48.125 | -2.886 | 51.875
960.2 | -3.214 | 48.144 | -2.891 | 51.856
960.641 | -3.213 | 48.162 | -2.894 | 51.838
961.082 | -3.211 | 48.187 | -2.896 | 51.813
961.523 | -3.207 | 48.22 | -2.898 | 51.78
961.964 | -3.201 | 48.257 | -2.898 | 51.743
962.405 | -3.195 | 48.29 | -2.898 | 51.71
962.846 | -3.193 | 48.31 | -2.899 | 51.69
963.287 | -3.193 | 48.321 | -2.901 | 51.679
963.727 | -3.194 | 48.327 | -2.903 | 51.673
964.168 | -3.193 | 48.339 | -2.904 | 51.661
964.609 | -3.189 | 48.358 | -2.904 | 51.642
965.05 | -3.187 | 48.369 | -2.903 | 51.631
965.491 | -3.188 | 48.37 | -2.904 | 51.63
965.932 | -3.189 | 48.372 | -2.906 | 51.628
966.373 | -3.189 | 48.379 | -2.907 | 51.621
966.814 | -3.185 | 48.387 | -2.905 | 51.613
967.255 | -3.181 | 48.404 | -2.904 | 51.596
967.695 | -3.179 | 48.438 | -2.907 | 51.562
968.136 | -3.177 | 48.479 | -2.913 | 51.521
968.577 | -3.176 | 48.513 | -2.917 | 51.487
969.018 | -3.174 | 48.521 | -2.917 | 51.479
969.459 | -3.172 | 48.513 | -2.913 | 51.487
969.9 | -3.171 | 48.53 | -2.915 | 51.47
970.341 | -3.169 | 48.572 | -2.921 | 51.428
970.782 | -3.164 | 48.625 | -2.925 | 51.375
971.222 | -3.159 | 48.67 | -2.928 | 51.33
971.663 | -3.159 | 48.694 | -2.932 | 51.306
972.104 | -3.16 | 48.713 | -2.937 | 51.287
972.545 | -3.16 | 48.729 | -2.939 | 51.271
972.986 | -3.159 | 48.719 | -2.936 | 51.281
973.427 | -3.158 | 48.699 | -2.932 | 51.301
973.868 | -3.157 | 48.72 | -2.934 | 51.28
974.309 | -3.155 | 48.757 | -2.938 | 51.243
974.75 | -3.154 | 48.759 | -2.938 | 51.241
975.19 | -3.153 | 48.755 | -2.937 | 51.245
975.631 | -3.15 | 48.773 | -2.937 | 51.227
976.072 | -3.146 | 48.82 | -2.941 | 51.18
976.513 | -3.142 | 48.889 | -2.949 | 51.111
976.954 | -3.14 | 48.935 | -2.955 | 51.065
977.395 | -3.139 | 48.951 | -2.957 | 51.049
977.836 | -3.138 | 48.951 | -2.956 | 51.049
978.277 | -3.136 | 48.946 | -2.953 | 51.054
978.717 | -3.135 | 48.941 | -2.951 | 51.059
979.158 | -3.135 | 48.942 | -2.951 | 51.058
979.599 | -3.135 | 48.961 | -2.955 | 51.039
980.04 | -3.133 | 49.013 | -2.962 | 50.987
980.481 | -3.128 | 49.087 | -2.969 | 50.913
980.922 | -3.125 | 49.124 | -2.973 | 50.876
981.363 | -3.124 | 49.122 | -2.971 | 50.878
981.804 | -3.125 | 49.103 | -2.969 | 50.897
982.245 | -3.125 | 49.095 | -2.967 | 50.905
982.685 | -3.12 | 49.131 | -2.969 | 50.869
983.126 | -3.115 | 49.178 | -2.972 | 50.822
983.567 | -3.112 | 49.223 | -2.977 | 50.777
984.008 | -3.109 | 49.272 | -2.983 | 50.728
984.449 | -3.106 | 49.32 | -2.988 | 50.68
984.89 | -3.103 | 49.339 | -2.989 | 50.661
985.331 | -3.102 | 49.341 | -2.987 | 50.659
985.772 | -3.101 | 49.351 | -2.988 | 50.649
986.212 | -3.1 | 49.373 | -2.991 | 50.627
986.653 | -3.099 | 49.413 | -2.997 | 50.587
987.094 | -3.096 | 49.46 | -3.003 | 50.54
987.535 | -3.092 | 49.509 | -3.007 | 50.491
987.976 | -3.088 | 49.56 | -3.012 | 50.44
988.417 | -3.084 | 49.61 | -3.016 | 50.39
988.858 | -3.079 | 49.653 | -3.019 | 50.347
989.299 | -3.075 | 49.689 | -3.021 | 50.311
989.74 | -3.071 | 49.714 | -3.022 | 50.286
990.18 | -3.069 | 49.733 | -3.023 | 50.267
990.621 | -3.069 | 49.752 | -3.026 | 50.248
991.062 | -3.067 | 49.788 | -3.03 | 50.212
991.503 | -3.061 | 49.847 | -3.034 | 50.153
991.944 | -3.058 | 49.895 | -3.039 | 50.105
992.385 | -3.056 | 49.925 | -3.043 | 50.075
992.826 | -3.055 | 49.949 | -3.046 | 50.051
993.267 | -3.052 | 49.975 | -3.048 | 50.025
993.707 | -3.047 | 50.016 | -3.05 | 49.984
994.148 | -3.043 | 50.057 | -3.053 | 49.943
994.589 | -3.042 | 50.089 | -3.058 | 49.911
995.03 | -3.042 | 50.111 | -3.062 | 49.889
995.471 | -3.041 | 50.124 | -3.063 | 49.876
995.912 | -3.038 | 50.142 | -3.063 | 49.858
996.353 | -3.033 | 50.168 | -3.063 | 49.832
996.794 | -3.03 | 50.202 | -3.065 | 49.798
997.234 | -3.028 | 50.241 | -3.07 | 49.759
997.675 | -3.024 | 50.285 | -3.073 | 49.715
998.116 | -3.019 | 50.338 | -3.078 | 49.662
998.557 | -3.013 | 50.408 | -3.084 | 49.592
998.998 | -3.009 | 50.464 | -3.089 | 49.536
999.439 | -3.008 | 50.487 | -3.093 | 49.513
999.88 | -3.009 | 50.469 | -3.09 | 49.531
1000.321 | -3.011 | 50.429 | -3.085 | 49.571
1000.762 | -3.013 | 50.414 | -3.084 | 49.586
1001.202 | -3.014 | 50.419 | -3.087 | 49.581
1001.643 | -3.014 | 50.465 | -3.094 | 49.535
1002.084 | -3.012 | 50.522 | -3.102 | 49.478
1002.525 | -3.007 | 50.585 | -3.108 | 49.415
1002.966 | -3.002 | 50.636 | -3.113 | 49.364
1003.407 | -2.998 | 50.67 | -3.115 | 49.33
1003.848 | -2.997 | 50.685 | -3.116 | 49.315
1004.289 | -2.997 | 50.693 | -3.117 | 49.307
1004.729 | -2.993 | 50.72 | -3.118 | 49.28
1005.17 | -2.988 | 50.757 | -3.12 | 49.243
1005.611 | -2.984 | 50.798 | -3.123 | 49.202
1006.052 | -2.981 | 50.844 | -3.127 | 49.156
1006.493 | -2.977 | 50.901 | -3.133 | 49.099
1006.934 | -2.972 | 50.962 | -3.139 | 49.038
1007.375 | -2.966 | 51.025 | -3.144 | 48.975
1007.816 | -2.964 | 51.072 | -3.15 | 48.928
1008.257 | -2.964 | 51.109 | -3.157 | 48.891
1008.697 | -2.962 | 51.146 | -3.161 | 48.854
1009.138 | -2.959 | 51.186 | -3.165 | 48.814
1009.579 | -2.955 | 51.238 | -3.17 | 48.762
1010.02 | -2.95 | 51.289 | -3.174 | 48.711
1010.461 | -2.947 | 51.332 | -3.178 | 48.668
1010.902 | -2.944 | 51.36 | -3.18 | 48.64
1011.343 | -2.941 | 51.376 | -3.18 | 48.624
1011.784 | -2.937 | 51.406 | -3.181 | 48.594
1012.224 | -2.932 | 51.444 | -3.183 | 48.556
1012.665 | -2.928 | 51.477 | -3.185 | 48.523
1013.106 | -2.925 | 51.501 | -3.186 | 48.499
1013.547 | -2.923 | 51.502 | -3.184 | 48.498
1013.988 | -2.921 | 51.515 | -3.184 | 48.485
1014.429 | -2.917 | 51.556 | -3.188 | 48.444
1014.87 | -2.915 | 51.6 | -3.193 | 48.4
1015.311 | -2.913 | 51.645 | -3.199 | 48.355
1015.752 | -2.911 | 51.691 | -3.205 | 48.309
1016.192 | -2.908 | 51.737 | -3.209 | 48.263
1016.633 | -2.9 | 51.768 | -3.208 | 48.232
1017.074 | -2.894 | 51.799 | -3.206 | 48.201
1017.515 | -2.891 | 51.842 | -3.211 | 48.158
1017.956 | -2.887 | 51.886 | -3.215 | 48.114
1018.397 | -2.882 | 51.927 | -3.217 | 48.073
1018.838 | -2.881 | 51.96 | -3.221 | 48.04
1019.279 | -2.881 | 51.989 | -3.227 | 48.011
1019.719 | -2.878 | 52.036 | -3.232 | 47.964
1020.16 | -2.873 | 52.087 | -3.236 | 47.913
1020.601 | -2.869 | 52.106 | -3.235 | 47.894
1021.042 | -2.866 | 52.124 | -3.235 | 47.876
1021.483 | -2.863 | 52.149 | -3.237 | 47.851
1021.924 | -2.859 | 52.204 | -3.242 | 47.796
1022.365 | -2.854 | 52.295 | -3.253 | 47.705
1022.806 | -2.85 | 52.349 | -3.258 | 47.651
1023.247 | -2.847 | 52.373 | -3.259 | 47.627
1023.687 | -2.845 | 52.402 | -3.262 | 47.598
1024.128 | -2.842 | 52.437 | -3.266 | 47.563
1024.569 | -2.835 | 52.516 | -3.272 | 47.484
1025.01 | -2.828 | 52.592 | -3.278 | 47.408
1025.451 | -2.823 | 52.651 | -3.284 | 47.349
1025.892 | -2.818 | 52.706 | -3.288 | 47.294
1026.333 | -2.813 | 52.758 | -3.292 | 47.242
1026.774 | -2.809 | 52.794 | -3.295 | 47.206
1027.214 | -2.807 | 52.82 | -3.297 | 47.18
1027.655 | -2.807 | 52.852 | -3.303 | 47.148
1028.096 | -2.806 | 52.889 | -3.309 | 47.111
1028.537 | -2.799 | 52.952 | -3.313 | 47.048
1028.978 | -2.794 | 53.006 | -3.317 | 46.994
1029.419 | -2.792 | 53.032 | -3.319 | 46.968
1029.86 | -2.788 | 53.091 | -3.325 | 46.909
1030.301 | -2.78 | 53.188 | -3.334 | 46.812
1030.741 | -2.774 | 53.237 | -3.337 | 46.763
1031.182 | -2.769 | 53.261 | -3.336 | 46.739
1031.623 | -2.769 | 53.287 | -3.341 | 46.713
1032.064 | -2.77 | 53.318 | -3.347 | 46.682
1032.505 | -2.764 | 53.381 | -3.352 | 46.619
1032.946 | -2.757 | 53.45 | -3.357 | 46.55
1033.387 | -2.749 | 53.529 | -3.363 | 46.471
1033.828 | -2.746 | 53.564 | -3.366 | 46.436
1034.269 | -2.748 | 53.557 | -3.367 | 46.443
1034.709 | -2.747 | 53.589 | -3.372 | 46.411
1035.15 | -2.745 | 53.64 | -3.378 | 46.36
1035.591 | -2.735 | 53.705 | -3.379 | 46.295
1036.032 | -2.725 | 53.771 | -3.381 | 46.229
1036.473 | -2.726 | 53.833 | -3.393 | 46.167
1036.914 | -2.725 | 53.897 | -3.404 | 46.103
1037.355 | -2.722 | 53.964 | -3.412 | 46.036
1037.796 | -2.718 | 54.008 | -3.416 | 45.992
1038.236 | -2.714 | 54.032 | -3.416 | 45.968
1038.677 | -2.709 | 54.065 | -3.417 | 45.935
1039.118 | -2.704 | 54.102 | -3.418 | 45.898
1039.559 | -2.703 | 54.134 | -3.423 | 45.866
1040 | -2.702 | 54.169 | -3.427 | 45.831